This lot of product meets the accepted performance criteria recommended in the USP and in compliance with ISO 11138-1.

PRODUCT NAME:  Quality Control Suspension			 FOR USE IN MONITORING: Growth Capabilities

ORGANISM: Aspergillus brasiliensis, ATCC® 16404

|  | AB1-1 |  | 2021-01-01 |  | 2021-07-01 |
| --- | --- | --- | --- | --- | --- |
|  | UAB-01 | POPULATION: |  | ≤100 Colony Forming Units/0.1 mL |  |
| Number of Passages: 2 |  | VOLUME: |  | 10 mL each vial |  |

Instructions for Use

- Manually shake or vortex Vial of Suspension.

- Remove cap, insert pipette or insert syringe through cap septum and withdraw 0.1 mL of Suspension.

- Deposit Suspension in Soybean Casein Digest tube of broth.

- Repeat process in steps 1 – 3 until the target number of tubes is achieved.

- For a Positive Control, deposit 0.1 mL of Suspension in a previously accepted lot of media.

- An un-inoculated tube of media can serve as a Negative Control.

- Incubate tubes at 30-35ºC for 2 – 3 days.

- Evaluate tubes for growth as evidenced by grayish-white cottony or feathery growth.

- Interpretation of Results:

Growth of the organism in the inoculated tubes = Pass, media is capable of supporting growth

Growth in Positive Control = Pass

No Growth in the Negative Control = Pass

| Storage and Shelf Life |  |  |  |
| --- | --- | --- | --- |
|  | Refrigerate at 2°C to 8°C |  | Keep away from sunlight |
|  | Do not freeze |  | Protect from heat, radioactive sources, & sterilizing agents |
| Shelf-life | 6 Months from the date of manufacture. |  |  |
|  | Do not used damaged vials of Spore Suspension. Do not use after expiration date. Spore Suspensions contain live cultures and should be handled with care. |  |  |

Dispose:  Autoclave for not less than 30 minutes at 121ºC or per other validated disposal cycle prior to discard.

| _____________________________ | June 16, 2021 |
| --- | --- |
| QUALITY ASSURANCE APPROVAL | DATE |